## NEW FRM-33-12 -SAE_clean
| Adverse Events and Device Deficiency Sheet Study ID: | [6 empty] | H-34
|  |  |  | Role: |  | Date:
| Written by: Baldassi Domizia |  |  | AJCL
| Approved by: Federica Azzimonti |  |  | RMCL
Adverse Events | [41 empty] | COMMUNICATION TO PMS |  |  | PMS feedback |  |  |  |  | Medical Expert Feedback | [6 empty] | Final evaluation
Site ID | Patient ID | Id AE | Report type | Initial Report Date | Report Date | Date of Onset | Date of Study Team Aware | Time team became aware (24 hr clock) | Name of the medical device involved | Components involved | ID number of study subject | Age | Gender | Attribution within the clinical investigation | Adverse Event (diagnosis, if known, or signs/ symptoms) | Event narrative | Admission date (dd/mmm/yyyy) | Discharge date (dd/mmm/yyyy) | Is serious event : | Classification of the adverse event | Outcome of the event | End Date | Relevant concomitant drugs and date of administration | Detail all possible and suspected causes including relevant medical history | Current clinical status | Severity | Causality: relationship to study medical device | Causality: relationship to study procedure | Expectedness | Action taken | Detailed treatment | Was the device permanently removed? | Device removal date | Attachments: Relevant Medical History | Attachments: Relevant Medical Records | Attachments: Laboratory Results | Attachments: Imaging Methods (e.g., X-ray, CT-scan) | Attachments: Other | Submitter | Name of submitter | Notes | communicated to PMS | if Yes, date | Date of reply | PMS Feedback | Complaint Number | Device relationship for PMS dept | Relationship to Procedure for PMS Dept | PMS Final feedback | Ready for Medical Expert Review | To be sent to ME | Date of Communication to ME | Medical Expert Causal relationship with the device (according to MDCG 2020-10-1) | Medical Expert rational / Consideration in case of possible / Probable / Causal Relationship with Device | Medical Expert Causal relationship with procedure (according to MDCG 2020-10-1) | Medical Expert rational / Consideration in case of possible / Probable / Causal Relationship with Procedure | MDCG Final Attribution for Relationship to device Sponsor | MDCG Final Attribution for Relationship to Procedure Sponsor
Samodzielny Publiczny Szpital Kliniczny im. Prof. Adama Grucy CMKP | 10 | 905 | Final | 2022-01-19 | 2022-11-04 | 2022-01-14 | 2022-01-14 | 08:51 | N/A | N/A | 101 - 10 | 76 | Male | Treatment Group | Dyspnoea, fluid collection in lungs | Chronic heart failure exacerbation | 2021-12-29 23:00 | 2022-01-17 23:00 | Yes | Hospitalization or prolongation of hospitalization | Recovered | 2022-06-08 22:00 | Furosemid 40mg 2xday from 14.01.2022. | Chronic heart failure | Good | Moderate | Not Related | Not Related | N/A | Pharmacological | Farmacological treatment: Furosemid 40mg 2xday from 14.01.2022 to 18.01.2022. With this treatment we got an improvement. | No |  | None | None | None | None | None | Investigator | Pawel Bartosz
Samodzielny Publiczny Szpital Kliniczny im. Prof. Adama Grucy CMKP | 13 | 951 | Final | 2022-06-09 | 2022-11-04 | 2022-04-12 | 2022-04-12 | 00:00 | Modulus | Stem | 101 - 13 | 41 | Male | Treatment Group | Intraoperative fracture of great trochanter | During the reduction of endoprotsthesis Intraoperatively, it came to fracture of great trochanter . Without dispacement. It was fixed with two Dall- Miles loops. | 2022-04-10 22:00 | 2022-04-18 22:00 | No |  | Recovered | 2022-07-03 22:00 | - | - | Rehabilitation. | Mild | Not Related | Probable | Anticipated | Not pharmacological | Fracture was fixed with two Dall- Miles loops. | No |  | None | None | None | Available | None | Investigator | Julia Macias | [10 empty] | yes
Samodzielny Publiczny Szpital Kliniczny im. Prof. Adama Grucy CMKP | 14 | 954 | Follow-up | 2022-06-17 | 2022-11-04 | 2022-06-04 | 2022-06-05 | 08:00 | Delta Revision cup/Modulus-R | cup/stem | 101 - 14 | 40 | Female | Treatment Group | Dislocation of left hip | at night, when standing up of bed she felt pain | 2022-06-03 22:00 | 2022-06-12 22:00 | Yes | Medical or surgical intervention required to prevent any of the above | Ongoing |  | - | Suspected cause is initial migration of the cup. | Reoperation is planned. | Severe | Not Related | Possible | Anticipated | Not pharmacological | 13.06.2022- reposition of the hip; | No |  | None | None | None | Available | None | Investigator | Julia Macias |  | yes | 2022-06-17 | 2022-06-20 |  | 279-22 |  |  |  |  | no
Samodzielny Publiczny Szpital Kliniczny im. Prof. Adama Grucy CMKP | 5 | 919 | Final | 2022-02-03 | 2022-11-04 | 2021-10-20 | 2021-10-20 | 00:00 | Revision Hip /Anca | Stem17mm/200mm | 101 - 5 | 76 | Female | Treatment Group | Intraoperative fracture of femur shaft | The femur shaft was weakened beacuse of previous use of metal loop; | 2021-10-17 22:00 | 2021-10-26 22:00 | Yes | Hospitalization or prolongation of hospitalization | Recovered | 2021-11-21 23:00 |  | The femur shaft was weakened beacuse of previous use of metal loop; 25 years ago THA of both hips, 2011- revision of right hipEcofit; | On the first control visit 22 Nov 2021 patient is walking with 2 crutches; | Severe | Not Related | Not Related | N/A | Other intraoperative reduction and fixation of the fracture |  | No |  | None | None | None | None | None | Investigator | Julia Macias | [10 empty] | yes
Samodzielny Publiczny Szpital Kliniczny im. Prof. Adama Grucy CMKP | 9 | 869 | Final | 2021-12-07 | 2022-11-04 | 2021-12-01 | 2021-12-01 | 08:00 | Delta Revision Cup | Delta Revision Cup, Hemispherical module, Bone screws, insert | 101 - 9 | 60 | Male | Treatment Group | Cup loosening | Early postoperative cup loosening, | 2021-11-18 23:00 | 2021-12-09 23:00 | Yes | Medical or surgical intervention required to prevent any of the above | Recovered | 2021-12-02 23:00 | None | Probable cause is technical mistake at first operation with Delta Revision implantation. Too proximal cup implantation and not sufficient cup hook stabilization at the ischial bone. Second cause could be wrong position of stable stem, various position in femur. | Patient undergone revision surgery at 2021-12-03, with cup and stem reimplantation. | Severe | Not Related | Not Related | N/A | Not pharmacological | Revision surgery with cup and stem removal. Reimplantation Trident Multihole cup with augment and Restoration stem. | Yes | 03/12/2021 | None | None | None | Available | None | Investigator | Pawel Bartosz |  | yes | 2021-12-09 | 2021-12-10 |  | 01-22 | not related |  |  |  | no
| Adverse events |  |  |  |  | 5
| Seriousness | n | % | Seriousness type | n | %
| Serious | 4 | 80 | Death | 0 | 0
|  |  |  | Life-threatening | 0 | 0
|  |  |  | Hospitalization or prolongation of hospitalization | 2 | 50
|  |  |  | Permanent impairment | 0 | 0
|  |  |  | Medical or surgical intervention required to prevent any of the above | 2 | 50
| Not serious | 1 | 20
| Severity | n | %
| Mild | 1 | 20
| Moderate | 1 | 20
| Severe | 3 | 60
| Causality: relationship to study medical device | n | % | Causality: relationship to study procedure | n | %
| Causal relationship (related) | 0 | 0 | Causal relationship (related) | 0 | 0
| Probable | 0 | 0 | Probable | 1 | 20
| Possible | 0 | 0 | Possible | 1 | 20
| Not related | 5 | 100 | Not related | 3 | 60
| Action taken | n | % | Detailed action | n | %
| Not pharmacological | 3 | 60 | Fracture fixed with two Dall- Miles loops. | 1 | 20
|  |  |  | Reposition of the hip | 1 | 20
|  |  |  | Revision surgery with cup and stem removal. Reimplantation Trident Multihole cup with augment and Restoration stem. | 1 | 20
| Pharmacological | 1 | 20 | Furosemid 40mg 2xday | 1 | 20
| Other | 1 | 20 | Intraoperative reduction and fixation of the fracture | 1 | 20
| Outcome | n | %
| Recovered | 4 | 80
| Recovered with residual effect | 0 | 0
| Ongoing | 1 | 20
| Fatal | 0 | 0
| Unknown | 0 | 0

## FRM 33-12 - DEVICE DEFICIENCY
Adverse Events and Device Deficiency Sheet Study ID: H-34
Device Deficiency
Patient ID |  | Onset date | Name of medical device | Components | Classification of deficiency | Description of deficiency and action taken | Could the deficiency have led to a SAE? | Notes | Initial/Follow-up | Clinical Dept.Revision | Date of communication to PMS | Date of PMS reply | PMS feedback
| Role: |  |  |  |  | Date: |  |  | Signature:
| Written by: Domizia Baldassi |  |  |  | AJCL
| Approved by: Federica Azzimonti |  |  |  | RMCL