|  | CLINICAL INVESTIGATION REPORT |  |  |  |
| --- | --- | --- | --- | --- |
|  | Department | Rep. Nr | Date | Rev. |
|  | Clinical Research | 01 | 11-Dec-23 | 1.0 |

| Study Title | An open label, observational, prospective, longitudinal cohort study to evaluate safety, clinical and radiographic outcomes of total hip arthroplasty with DELTA Revision acetabular cup. |
| --- | --- |
| Short Title | DELTA REVISION CUP |
| Protocol ID | H-34 |
| Product | CE marked DELTA Revision acetabular cup |
| Study Phase | Enrollment |
| Subject Population | 49 |
| Sponsor | Limacorporate Spa Via Nazionale 52, 33030, Villanova di San Daniele (UD), Italy |
| Sponsor representative | Federica Azzimonti Clinical Research Manager, Limacorporate S.p.A. |
| Principal Investigator | Dr. Jerzy Bialecki Samodzielny Publiczny Szpital Kliniczny im. Prof. Adama Grucy CMKP in Otwock, Poland |
| Sub-Investigator(s) | Dr. Julia Macias Dr. Pawel Bartosz Samodzielny Publiczny Szpital Kliniczny im. Prof. Adama Grucy CMKP in Otwock, Poland |
| Statistical Analysis | Intermediate statistical analysis: Limacorporate S.p.A. |

Clinical investigation identified as H-34 and entitled “An open label, observational, prospective, longitudinal cohort study to evaluate safety, clinical and radiographic outcomes of total hip arthroplasty with DELTA

Revision acetabular cup” has been conducted in accordance with ISO 14155:2020 “Clinical investigation of medical devices for human subjects — Good Clinical Practice” and with the ethical principles of the Declaration of Helsinki.

Table of contents

List of Abbreviations	3

List of Tables	4

List of Figures	4

1.0	SUMMARY	5

2.0	INTRODUCTION	7

3.0	RESULTS	8

3.1 Study Flowchart	8

3.2 Patient population and preoperative data	9

3.3 Intraoperative data	11

3.3.1 Surgery data	11

3.3.2 Implant information	12

3.4 Efficacy analysis	13

3.4.1 Clinical results	13

3.4.2 Radiographic results	14

3.5 Safety analysis	18

3.5.1 Adverse events reporting	18

3.5.2 Device deficiency	20

3.5.3 Device survivorship	20

4.0	SUMMARY AND OUTLOOK	20

# List of Abbreviations

| ADE | Adverse Device Effect |
| --- | --- |
| AE | Adverse Event |
| BMI | Body-Mass Index |
| eCRF | electronic Case Report Form |
| FU | Follow-Up |
| HHS | Harris Hip Score |
| ISO | International Organization for Standardization |
| OHS | Oxford Hip Score |
| ROM | Range of Motion |
| SADE | Serious Adverse Device Effect |
| SAE | Serious Adverse Event |
| SD | Standard Deviation |
| THA | Total Hip Arthroplasty |
| UHMWPE | Ultra High Molecular Weight Polyethylene |
| USADE | Unexpected Serious Adverse Device Effect |

# List of Tables

Table 1 Demographics of study population.	9

Table 2 Gender, working status, activity level, smoking and drinking habits, concomitant medications of study population.	9

Table 3 Preoperative assessment of study population.	10

Table 4 Surgical data of study population with regard to surgical approach, anesthesia, intraoperative complications and prophylaxis therapy.	11

Table 5 Surgery data of study population in terms of surgery time and hematocrit (intraoperative and postoperative).	11

Table 6 bone stock quality and bone grafting on the acetabular side of study population	11

Table 7 Bone stock quality and bone grafting on the femoral side of study population	12

Table 8 Implant details of the acetabular side of study population	12

Table 9 Implant details of the femoral head of study population	13

Table 10 Implant used on the femoral side of study population.	13

Table 11 Change in HHS from preoperative (baseline) to 2-year follow-up	13

Table 12 Change in total score description of HHS from preoperative (baseline) to 2-year follow-up.	14

Table 13 Change in OHS from preoperative (baseline) to 2-year follow-up.	14

Table 14 Change in total score description of OHS from preoperative (baseline) to 2-year follow-up.	14

Table 15 Preoperative radiografic assessment	15

Table 16 Postoperative assessment of acetabular radiographic migration	15

Table 17 Postoperative assessment of radiolucent lines on the acetabular cup	16

Table 18 Postoperative assessment of osteolysis on the acetabular cup	16

Table 19 Postoperative radiographic assessment (presence of distant femoral component migration/subsidence, presence of periprosthetic fractures, dislocation, loosening, bone remodeling cortical hypertrophy/stress shielding).	17

Table 20 Postoperative radiographic assessment (heterotopic ossification, infection, presence of hardware failure, cysts, osteophytes, sclerosis/calcifications).	18

Table 21 Adverse events occurred (seriousness, severity, causal relationship with study medical device and study procedure, outcome of the event).	18

Table 22 Detailed description of adverse events occurred.	19

# List of Figures

Figure 1 Study flowchart	8

# SUMMARY

| Name of sponsor/Company | Limacorporate S.p.A |
| --- | --- |
| Study ID | H-34 |
| Study title | An open label, observational, prospective, longitudinal cohort study to evaluate safety, clinical and radiographic outcomes of total hip arthroplasty with DELTA Revision acetabular cup. |
| Name of study product | CE marked DELTA Revision acetabular cup |
| Investigators & study site(s) | Dr. Jerzy Bialecki Samodzielny Publiczny Szpital Kliniczny im. Prof. Adama Grucy CMKP in Otwock, Poland |
| Publication (reference) | N/A |
| Study period: | Date of first enrolment: 14/09/2021 Date of last enrolment: N/A Current phase: enrollment |
| Objectives: | Primary objective - To assess the clinical outcomes of the DELTA Revision acetabular cup used in real life settings over a period of 2 years Secondary objectives - To assess the radiographical and clinical outcomes of the DELTA Revision acetabular cup used in real life settings over a period of 2 years -To evaluate the safety profile of the DELTA Revision acetabular cup used in real life settings over a period of 5 years - To describe baseline data (demographic data, primary diagnosis, aetiology, affected side, type of surgery and intra-operative data) of patients with Delta Revision acetabular cup |
| Criteria for evaluation/Endpoints: | Primary endpoint: - The evaluation of Harris Hip Score (HHS) at pre-operative visit, intraoperatively/discharge, 2 months, 6 months, 1 year and 2 years after the implant. Secondary endpoints: - Evaluation of Range of Motion (ROM) at pre-operative visit, intraoperatively/discharge, 2 months, 6 months, 1 year and 2 years after the implant; - Evaluation of Oxford Hip Score (OHS) at pre-operative visit, intraoperatively/discharge, 2 months, 6 months, 1 year, 2 years, 3 years, 4 years and 5 years after the implant; - Survivorship of the implant (Kaplan-Meier estimate) at 2 years of follow-up after surgery; - Radiographic implant evaluation and stability assessment of the DELTA Revision acetabular cup at the following timepoints: at discharge, 2 months, 6 months, 1 year, and 2 years after the implant - Incidence, type and severity of all the Adverse Events (AEs), Serious Adverse Events (SAEs), Adverse Device Effects (ADEs) and Serious Adverse Device Effects (SADEs) at intraoperatively, discharge, 2 months, 6 months, 1 year, 2 years, 3 years, 4 years and 5 years after the implant. |
| Number of patients: | Planned sample size: 49 Enrolled: 21 Completed: 0 |
| Inclusion and exclusion criteria: | Inclusion criteria: Patients must meet the following criteria for study entry: - Male or female. - Age ≥ 18 years old. - All patients must give written informed consent approved by the study site’s Institutional Review Board (IRB)/Ethical Committee (EC). - Adult patients in whom a decision has already been made to perform a total hip arthroplasty with DELTA Revision acetabular cup as per indication for use. The decision to implant DELTA Revision acetabular cup must be taken prior to, and independently from the decision to enroll the patient. This decision should be made in accordance with routine clinical practice at the study site concerned. - Patient is able to comply with the protocol Exclusion criteria: A patient will not be included in the study if they meet any of the following criteria: - Adult patients with any DELTA Revision acetabular cup contraindication for use as reported in the current local Instruction for use. - For female patients, current pregnancy and/or lactation or planning a pregnancy. |
| Statistical methods: | An internal statistical analysis was performed on all collected data, using Microsoft® Excel. |
| Date of the report: | 11 December 2023 |

# INTRODUCTION

H-34 is a post-marketing, prospective, observational, monocentric, longitudinal cohort clinical study that aims at collecting clinical, patient-reported and radiographic early outcomes after total hip arthroplasty (THA) with DELTA Revision acetabular cup. The study is being carried out at the Samodzielny Publiczny Szpital Kliniczny im. Prof. Adama Grucy CMKP in Otwock, Poland. Data are being collected prospectively and following hospital standard of care. Every patient is followed from preoperative visit up to 5 years after surgery.

Study population consist of 49 patients in need for a hip revision surgery and deemed suitable for receiving a DELTA Revision acetabular cup. The expected enrollment period is 30 months. So far, 23 patients have been screened and 21 patients have been enrolled, of which 5 performed the 2-year follow-up visit.

Study visits are performed according to the local standard of care. Timelines and temporal windows are the following:

- Pre-operative visit

- Intraoperative visit

- Discharge visit

- Follow-Up 2 months: ± 30 days

- Follow-Up 6 months: ± 60 days

- Follow-Up 12 months: ± 6 months

- Follow-Up 24 months: ± 6 months

- Survival Follow-Up visit at 3,4 and 5 years: ± 6 months

Relevant data will be captured on an eCRF. These data will be transmitted to the sponsor for analysis, will be anonymous and will be identified by a patient number. As this study is non-interventional, all data collected will be in accordance with the practice of clinicians implanting DELTA Revision acetabular cup.

At pre-operative, intra-operative, discharge and each follow-up visit, the following information will be collected:

- Medical history and demographic data (only at pre-operative visit);

- Primary diagnosis and physical examination (only at pre-operative visit);

- Surgical data (only at intra-operative visit);

- X-rays as per standard care;

- Radiographic evaluation;

- Clinical assessment (HHS, ROM, OHS);

- Assessment of type, nature and severity of any DD, AE, ADE, SAE and SADE;

- Survivorship of implant.

The study is still in the enrollment phase and this intermediate clinical study report is based on eCRF data extraction performed on the 11th of December 2023.

# RESULTS

## 3.1 Study Flowchart

Figure 1 Study flowchart

## 3.2 Patient population and preoperative data

| Variable | N | Mean | SD | Min | Max | Median |
| --- | --- | --- | --- | --- | --- | --- |
| Pt enrolled | 21 |  |  |  |  |  |
| Weight (kg) | 21 | 78 | 14 | 57 | 110 | 79 |
| Height (cm) | 21 | 168 | 10 | 152 | 189 | 167 |
| BMI (kg/m2) | 21 | 28 | 4 | 22 | 35 | 27.3 |
| Age (years) | 21 | 64 | 13 | 40 | 87 | 68 |

Table 1 Demographics of study population.

| Variable | n | % |
| --- | --- | --- |
| Gender | 21 |  |
| Female | 12 | 57 |
| Male | 9 | 43 |
| Race | 21 |  |
| White | 21 | 100 |
| Working status | 21 |  |
| Retired | 10 | 48 |
| Not Working | 8 | 38 |
| Active | 2 | 10 |
| Missing data | 1 | 5 |
| Activity level | 21 |  |
| Sedentary | 17 | 81 |
| Normal | 3 | 14 |
| Intense | 0 | 0 |
| Missing Data | 1 | 5 |
| Smoking habits | 21 |  |
| No smoking | 16 | 76 |
| Smoking, daily | 5 | 24 |
| Alcohol drinking | 21 |  |
| No alcohol | 20 | 95 |
| Alcohol drinking, occasionally | 1 |  |
| Alcohol drinking, daily |  |  |
| Concomitant medication | 21 |  |
| Yes | 15 | 71 |
| No | 5 |  |
| - | 1 |  |

Table 2 Gender, working status, activity level, smoking and drinking habits, concomitant medications of study population.

| Variable | n | % |
| --- | --- | --- |
| Affected side | 21 |  |
| Left | 11 | 52.4 |
| Right | 10 | 47.6 |
| Primary diagnosis | 21 |  |
| Presence of bone stock of poor quality or inadequate for other reconstructive techniques as indicated by deficiencies of the acetabulum | 10 | 47.6 |
| Revision of previous unsuccessful femoral head replacement, cup arthroplasty or other procedure | 9 | 42.9 |
| Clinical management problem where arthrodesis or alternative reconstruction techniques are less likely to achieve satisfactory results | 2 | 9.5 |
| Previous hip surgeries on the affected side | 21 |  |
| Yes | 21 | 100 |
| No | 0 | 0 |
| Previous hip treatments or surgeries on the contralateral side | 21 |  |
| Yes | 8 | 38.1 |
| No | 13 | 61.9 |
| Description of pain | 21 |  |
| No pain | 3 | 14.3 |
| Unable to describe | 0 | 0 |
| At rest | 4 | 19 |
| Occasional start-up pain | 4 | 19 |
| Always start-up pain | 6 | 28.6 |
| With every movement | 3 | 14.3 |
| - | 1 | 4.8 |
| Pain therapy | 21 |  |
| None | 7 | 33.3 |
| NSAIDs, occasionally | 3 | 14.3 |
| NSAIDs, regularly | 4 | 19 |
| Steroids, occasionally | 0 | 0 |
| Steroids, regularly | 0 | 0 |
| Opiates, occasionally | 3 | 14.3 |
| Opiates, regularly | 4 | 19 |
| Osteoporosis | 21 |  |
| Yes | 0 | 0 |
| No | 20 | 95.2 |
| - | 1 | 4.8 |

Table 3 Preoperative assessment of study population.

## 3.3 Intraoperative data

### 3.3.1 Surgery data

| Variable | n | % |
| --- | --- | --- |
| Surgical approach | 21 |  |
| Postero-lateral | 21 | 100 |
| Anesthesia | 21 |  |
| Subarachnoid | 19 | 90.5 |
| General | 2 | 9.5 |
| Intraoperative complications | 21 |  |
| No | 19 | 90.5 |
| Yes, not implant-related | 2 | 9.5 |
| Antibiotic prophylaxis | 21 |  |
| Yes | 21 | 100 |
| No | 0 | 0 |
| Antithrombotic prophylaxis | 21 |  |
| Yes | 21 | 100 |
| No | 0 | 0 |
| Antihemorrhagic prophylaxis | 21 |  |
| Yes | 12 | 57.1 |
| No | 8 | 38.1 |

Table 4 Surgical data of study population with regard to surgical approach, anesthesia, intraoperative complications and prophylaxis therapy.

| Variable | n | Mean | SD | Min | Max | Median |
| --- | --- | --- | --- | --- | --- | --- |
| Surgery time | 21 | 158 | 37 | 105 | 235 | 155 |
| Intraoperative hematocrit | 21 | 41 | 4 | 32 | 50 | 41 |
| Immediate postoperative hematocrit | 21 | 31 | 5 | 23 | 39 | 31 |

Table 5 Surgery data of study population in terms of surgery time and hematocrit (intraoperative and postoperative).

| Variable | n | % |
| --- | --- | --- |
| Bone stock quality - acetabulum | 21 |  |
| Normal | 2 | 9.5 |
| Porotic | 17 | 81 |
| Sclerotic | 2 | 9.5 |
| Bone grafting | 21 |  |
| Yes | 16 | 76.2 |
| No | 5 | 23.8 |
| Graft material | 21 |  |
| Allograft chips | 8 | 38.1 |
| Autograft | 1 | 4.8 |
| Other | 8 | 38.1 |
| - | 4 | 19 |

Table 6 bone stock quality and bone grafting on the acetabular side of study population.

| Variable | n | % |
| --- | --- | --- |
| Bone stock quality - femur | 21 |  |
| Normal | 10 | 47.6 |
| Porotic | 10 | 47.6 |
| Sclerotic | 1 | 4.8 |
| Bone grafting | 21 |  |
| Yes | 1 | 4.8 |
| No | 20 | 95.2 |
| Graft material | 21 |  |
| - | 20 | 95.2 |
| Other | 1 | 4.8 |

Table 7 Bone stock quality and bone grafting on the femoral side of study population.

### 3.3.2 Implant information

| Variable | n | % | Variable | n | % |
| --- | --- | --- | --- | --- | --- |
| Cup type | 21 |  | Cup liner material | 21 |  |
| DELTA-REVISION | 14 | 66.7 | Neutral UHMWPE X-LIMA + Ti6Al4V | 2 | 9.5 |
| DELTA-REVISION+TT HEMISPHERIC MODULE ECC. 12 | 5 | 23.8 | Neutral UHMWPE X-LIMA + Ti6Al4V w/ angled 10° spacer | 1 | 4.8 |
| DELTA-REVISION+TT HEMISPHERIC MODULE ECC. 18 | 1 | 4.8 | Neutral UHMWPE X-LIMA + Ti6Al4V w/ angled 20° spacer | 1 | 4.8 |
| - | 1 | 4.8 | Protruded UHMWPE X-LIMA + Ti6Al4V | 7 | 33.3 |
| Cup diameter | 21 |  | Protruded UHMWPE X-LIMA + Ti6Al4V w/ neutral spacer | 3 | 14.3 |
| 50 mm | 1 | 4.8 | Protruded UHMWPE X-LIMA + Ti6Al4V w/ neutral +5 spacer | 1 | 4.8 |
| 54 mm | 4 | 19 | Protruded UHMWPE X-LIMA + Ti6Al4V w/ angled 20° spacer | 4 | 19 |
| 58 mm | 6 | 28.6 | - | 2 | 9.5 |
| 62 mm | 5 | 23.8 | Cup liner size | 21 |  |
| 66 mm | 4 | 19 | ID 32 mm-MEDIUM | 10 | 47.6 |
| - | 1 | 4.8 | ID 32 mm-LARGE | 1 | 4.8 |
| Cup cement | 21 |  | ID 36 mm-LARGE | 8 | 38.1 |
| No | 21 |  | - | 2 | 9.5 |

Table 8 Implant details of the acetabular side of study population.

| Variable | n | % |
| --- | --- | --- |
| Head type | 21 |  |
| Modular | 19 | 90.5 |
| - | 2 | 9.5 |
| Head material | 21 |  |
| Biolox (r) Delta | 15 | 71.4 |
| Biolox (r) Delta + Ti6AI4V | 1 | 4.8 |
| CoCrMo Taper 12/14 | 2 | 9.5 |
| CoCrMo Taper 14/16 | 1 | 4.8 |
| - | 2 | 9.5 |
| Head diameter | 21 | 100 |
| 32 mm | 11 | 52.4 |
| 36 mm | 8 | 38.1 |
| - | 2 | 9.5 |
| Head size | 21 |  |
| S | 6 | 28.6 |
| M | 4 | 19 |
| L | 7 | 33.3 |
| XL | 2 | 9.5 |
| - | 2 | 9.5 |

Table 9 Implant details of the femoral head of study population.

| Variable | n | % |
| --- | --- | --- |
| Stem type | 21 |  |
| H-MAX S standard | 2 | 9.5 |
| Modulus-R | 2 | 9.5 |
| Modulus Taper B | 3 | 14.3 |
| Revision | 4 | 19 |
| Competitor (no Lima) | 10 | 47.6 |

Table 10 Implant used on the femoral side of study population.

## 3.4 Efficacy analysis

### 3.4.1 Clinical results

#### 3.4.1.1 Clinical results – Harris Hip Score (HHS)

| Time point | n | Mean | SD | Min | Max | Median |
| --- | --- | --- | --- | --- | --- | --- |
| Preoperative | 21 | 38.6 | 15.7 | 12 | 64 | 41 |
| 2 months FU | 17 | 60.6 | 12.7 | 42 | 80 | 56 |
| 6 months FU | 14 | 73.9 | 15.5 | 45 | 94 | 74.5 |
| 1 year FU | 9 | 84.3 | 12.7 | 65 | 97 | 89 |
| 2 years FU | 5 | 77 | 13.8 | 57 | 93 | 79 |

Table 11 Change in HHS from preoperative (baseline) to 2-year follow-up

|  | Preoperative |  | 2 months FU |  | 6 months FU |  | 1 year FU |  | 2 years FU |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total score description | n | % | n | % | n | % | n | % | n | % |
| Very poor | 20 | 95.2 | 9 | 52,9 | 3 | 21.4 | 0 | 0 | 1 | 20 |
| Poor | 1 | 4.8 | 1 | 5,9 | 1 | 7.1 | 2 | 22.2 | 0 | 0 |
| Fair | 0 | 0 | 6 | 35,3 | 5 | 35.7 | 1 | 11.1 | 2 | 40 |
| Good | 0 | 0 | 1 | 5,9 | 1 | 7.1 | 2 | 22.2 | 1 | 20 |
| Excellent | 0 | 0 | 0 | 0,0 | 4 | 28,6. | 4 | 44.4 | 1 | 20 |

Table 12 Change in total score description of HHS from preoperative (baseline) to 2-year follow-up.

#### 3.4.1.2 Clinical results – Oxford Hip Score (OHS)

| Time point | n | Mean | SD | Min | Max | Median |
| --- | --- | --- | --- | --- | --- | --- |
| Preoperative | 21 | 19.7 | 8.9 | 3 | 37 | 21 |
| 2 months FU | 17 | 30.9 | 7.6 | 17 | 43 | 33 |
| 6 months FU | 14 | 38.6 | 5.8 | 28 | 47 | 38.5 |
| 1 year FU | 9 | 41.2 | 6.3 | 28 | 48 | 42 |
| 2 years FU | 5 | 39.4 | 8.5 | 25 | 46 | 42 |

Table 13 Change in OHS from preoperative (baseline) to 2-year follow-up.

|  | Preoperative |  | 2 months FU |  | 6 months FU |  | 1 year FU |  | 2 years FU |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Total score description | n | % | n | % | n | % | n | % | n | % |
| Very poor | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0. | 0 | 0 |
| Poor | 17 | 81 | 7 | 41.2 | 0 | 0 | 0 | 0. | 1 | 20 |
| Fair | 2 | 9.5 | 2 | 11.8 | 3 | 2.4 | 1 | 11.1 | 0 | 0 |
| Good | 2 | 9.5 | 6 | 35.3 | 6 | 42.9 | 2 | 22.2 | 1 | 20 |
| Excellent | 0 | 0 | 2 | 11.8 | 5 | 35.7 | 6 | 66.7 | 3 | 60 |

Table 14 Change in total score description of OHS from preoperative (baseline) to 2-year follow-up.

### 3.4.2 Radiographic results

#### 3.4.2.1 Preoperative radiographic evaluation

| Variable |  | n | % |  |
| --- | --- | --- | --- | --- |
| Varus/valgus deformity |  | 21 |  |  |
| Neutral |  | 19 | 90.5 |  |
| Varus |  | 2 | 9.5 |  |
| Valgus |  | 0 | 0 |  |
| Osteoarthritis severity |  | 21 |  |  |
| Grade I |  | 11 | 52.4 |  |
| Grade II |  | 2 | 9.5 |  |
| Grade III |  | 0 | 0 |  |
| Grade IV |  | 4 | 19 |  |
| - |  | 4 | 19 |  |
| Presence of osteophytes |  | 21 |  |  |
| Yes |  | 1 | 4.8 |  |
| No |  | 17 | 81 |  |
| Missing Data |  | 3 | 14.3 |  |
| Presence of cysts | 21 |  |  |  |
| Yes | 1 |  |  | 4.8 |
| No | 17 |  |  | 81 |
| Missing Data | 3 |  |  | 14.3 |
| Sclerosis | 21 |  |  |  |
| None | 15 |  |  | 71.4 |
| Slight | 4 |  |  | 19 |
| Moderate | 2 |  |  | 9.5 |

Table 15 Preoperative radiografic assessment

#### 3.4.2.2 Postoperative radiographic evaluation

|  | Immediate postoperative |  |  | 2 months FU |  | 6 months FU |  | 12 months FU |  | 24 months FU |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Variable | n | % | n |  | % | n | % | n | % | n | % |
| Acetabular component migration | 20 |  | 17 |  |  | 14 |  | 9 |  | 5 |  |
| Absent | 19 | 95 | 17 |  | 100 | 14 | 100 | 9 | 100 | 5 | 100 |
| Present | 1 | 5 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |

Table 16 Postoperative assessment of acetabular radiographic migration

|  |  | Zone 1 |  | Zone 2 |  | Zone 3 |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Immediate postoperative | Radiolucent lines | n | % | n | % | n | % |
|  | None | 16 | 80 | 15 | 75 | 19 | 95 |
|  | <1 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | ≥1 to <2 mm | 0 | 0 | 2 | 10 | 0 | 0 |
|  | ≥ 2 to < 3 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | ≥3 to <4 mm | 0 | 0 | 1 | 5 | 0 | 0 |
|  | ≥ 4 mm | 4 | 20 | 2 | 10 | 1 | 5 |
| 2 months FU | Radiolucent lines | n | % | n | % | n | % |
|  | None | 11 | 65 | 14 | 82 | 16 | 94 |
|  | <1 mm | 0 | 0 | 0 | 0 | 1 | 6 |
|  | ≥1 to <2 mm | 0 | 0 | 2 | 12 | 0 | 0 |
|  | ≥ 2 to < 3 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | ≥3 to <4 mm | 2 | 12 | 0 | 0 | 0 | 0 |
|  | ≥ 4 mm | 4 | 24 | 1 | 6 | 0 | 0 |
| 6 months FU | Radiolucent lines | n | % | n | % | n | % |
|  | None | 14 | 100 | 9 | 64 | 13 | 93 |
|  | <1 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | ≥1 to <2 mm | 0 | 0 | 3 | 21 | 0 | 0 |
|  | ≥2 to <3 mm | 0 | 0 | 1 | 7 | 1 | 7 |
|  | ≥3 to <4 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | ≥ 4 mm | 0 | 0 | 1 | 7 | 0 | 0 |
| 1 year FU | Radiolucent lines | n | % | n | % | n | % |
|  | None | 8 | 89 | 7 | 78 | 9 | 100 |
|  | <1 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | ≥1 to <2 mm | 1 | 11 | 1 | 11 | 0 | 0 |
|  | ≥2 to <3 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | ≥3 to <4 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | ≥ 4 mm | 0 | 0 | 1 | 11 | 0 | 0 |
| 2 years FU | Radiolucent lines | n | % | n | % | n | % |
|  | None | 5 | 100 | 4 | 80 | 4 | 80 |
|  | <1 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | ≥1 to <2 mm | 0 | 0 | 0 | 0 | 1 | 20 |
|  | ≥2 to <3 mm | 0 | 0 | 1 | 20 | 0 | 0 |
|  | ≥3 to <4 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | ≥ 4 mm | 0 | 0 | 0 | 0 | 0 | 0 |

Table 17 Postoperative assessment of radiolucent lines on the acetabular cup

|  |  | Zone 1 |  | Zone 2 |  | Zone 3 |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Immediate postoperative | Osteolysis acetabular component | n | % | n | % | n | % |
|  | None | 18 | 90 | 18 | 90 | 19 | 95 |
|  | 1 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | 2 mm | 0 | 0 | 1 | 5 | 0 | 0 |
|  | > 2mm | 2 | 10 | 1 | 5 | 0 | 0 |
| 2 months FU | Osteolysis acetabular component | n | % | n | % | n | % |
|  | None | 13 | 76 | 16 | 94 | 17 | 100 |
|  | 1 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | 2 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | > 2mm | 4 | 24 | 1 | 6 | 0 | 0 |
| 6 months FU | Osteolysis acetabular component | n | % | n | % | n | % |
|  | None | 13 | 93 | 11 | 79 | 13 | 93 |
|  | 1 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | 2 mm | 1 | 7 | 2 | 14 | 0 | 0 |
|  | > 2mm | 0 | 0 | 1 | 7 | 1 | 7 |
| 1 year FU | Osteolysis acetabular component | n | % | n | % | n | % |
|  | None | 8 | 89 | 8 | 89 | 9 | 100 |
|  | 1 mm | 1 | 11 | 1 | 11 | 0 | 0 |
|  | 2 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | > 2mm | 0 | 0 | 0 | 0 | 0 | 0 |
| 2 years FU | Osteolysis acetabular component | n | % | n | % | n | % |
|  | None | 5 | 100 | 5 | 100 | 5 | 100 |
|  | 1 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | 2 mm | 0 | 0 | 0 | 0 | 0 | 0 |
|  | > 2mm | 0 | 0 | 0 | 0 | 0 | 0 |

Table 18 Postoperative assessment of osteolysis on the acetabular cup

|  | Immediate postoperative |  | 2 months FU |  |  | 6 months FU |  | 1 year FU |  | 2 years FU |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Variable | n | % | n |  | % | n | % | n | % | n | % |
| Presence of Distal Femoral Component Migration/Subsidence | 20 |  | 17 |  |  | 14 |  | 9 |  | 5 |  |
| Absent (<5 mm) | 20 | 100 | 17 |  | 100 | 14 | 100 | 9 | 100 | 5 | 100 |
| Present (>5 mm) | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Presence of any Periprosthetic Fracture | 20 |  | 17 |  |  | 14 |  | 9 |  | 5 |  |
| Absent | 17 | 85 | 17 |  | 100 | 14 | 100 | 9 | 100 | 5 | 100 |
| Present | 3 | 15 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Type of periprosthetic fractur | 3 |  | - |  |  | - |  | - |  | - |  |
| Type A | 2 | 66,7 |  |  |  |  |  |  |  |  |  |
| Type B1 | 1 | 33,3 |  |  |  |  |  |  |  |  |  |
| Dislocation | 20 |  | 17 |  |  | 14 |  | 9 |  | 5 |  |
| Absent | 20 | 100 | 17 |  | 100 | 14 | 100 | 9 | 100 | 5 | 100 |
| Loosening | 20 |  | 17 |  |  | 14 |  | 9 |  | 5 |  |
| Absent | 20 | 100 | 17 |  | 100 | 14 | 100 | 9 | 100 | 5 | 100 |
| Bone remodellingCortical Hypertrophy | 20 |  | 17 |  |  | 14 |  | 9 |  | 5 |  |
| Absent | 20 | 100 | 17 |  | 100 | 13 | 92.9 | 9 | 100 | 5 | 100 |
| Present | 0 | 0 | 0 |  | 0 | 1* | 7.1 | 0 | 0 | 0 | 0 |
| Bone remodelingStress Shielding | 20 |  | 17 |  |  | 13 |  | 9 |  | 5 |  |
| Absent | 20 | 100 | 17 |  | 100 | 13 | 100 | 9 | 100 | 5 | 100 |
| Present | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
|  |  |  |  |  |  | *New bone appear in the zone 2 of acetabulum |  |  |  |  |  |

Table 19 Postoperative radiographic assessment (presence of distant femoral component migration/subsidence, presence of periprosthetic fractures, dislocation, loosening, bone remodeling cortical hypertrophy/stress shielding).

|  | Immediate postoperative |  |  | 2 months FU |  | 6 months FU |  |  |  | 1 year FU |  |  | 2 years FU |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Variable | n | % | n |  | % | n | % |  |  | n | % |  | n | % |
| Heterotopic Ossification | 20 |  | 17 |  |  | 14 |  |  |  | 9 |  |  | 5 |  |
| Grade 0 | 18 | 90 | 12 |  | 70.6 | 9 | 64.3 |  |  | 6 | 66.7 |  | 2 | 40 |
| Grade 1 | 2 | 10 | 5 |  | 29.4 | 5 | 35.7 |  |  | 2 | 22.2 |  | 2 | 40 |
| Grade 2 | 0 | 0 | 0 |  | 0 | 0 | 0 |  |  | 1 | 11.1 |  | 1 | 20 |
| Infection | 20 |  | 17 |  |  | 14 |  |  |  | 9 |  |  | 5 |  |
| Absent | 20 | 100 | 17 |  | 100 | 14 | 100 |  |  | 9 | 100 |  | 5 | 100 |
| Presence of Hardware Failure | 20 |  | 17 |  |  | 14 |  |  |  | 9 |  |  | 5 |  |
| Absent | 20 | 100 | 0 |  | 0 | 14 | 100 |  |  | 9 | 100 |  | 5 | 100 |
| Presence of cysts | 20 |  | 17 |  |  | 14 |  |  |  | 9 |  |  | 5 |  |
| Absent | 20 | 100 | 17 |  | 100 | 14 | 100 |  |  | 9 | 100 |  | 5 | 100 |
| Presence of osteophytes | 20 |  | 16 |  |  | 14 |  |  |  | 9 |  |  | 5 |  |
| Absent | 20 | 100 | 16 |  | 100 | 14 |  | 100 |  | 9 |  | 100 | 5 | 100 |
| Presence of Sclerosis Calcifications | 20 |  | 17 |  |  | 14 |  |  |  | 9 |  |  | 5 |  |
| Absent | 20 | 100 | 17 |  | 100 | 13 |  | 92.9 |  | 9 |  | 100 | 5 | 100 |
| Present | 0 | 0 | 0 |  | 0 | 1 |  | 7.1 |  | 0 |  | 0 | 0 | 0 |

Table 20 Postoperative radiographic assessment (heterotopic ossification, infection, presence of hardware failure, cysts, osteophytes, sclerosis/calcifications).

## 3.5 Safety analysis

### 3.5.1 Adverse events reporting

| Adverse events | n | % | Seriousness type | n | % |
| --- | --- | --- | --- | --- | --- |
| Total adverse events | 8 |  | Death | 0 | 0 |
| Serious | 6 | 75 | Life-threatening | 0 | 0 |
|  |  |  | Hospitalization or prolongation of hospitalization | 4 | 66.7 |
|  |  |  | Permanent impairment | 0 | 0 |
|  |  |  | Medical or surgical intervention required to prevent any of the above | 2 | 33.3 |
| Not serious | 2 | 25 |  |  |  |
| Severity | n | % | Outcome of the event | n | % |
| Mild | 1 | 12.5 | Recovered | 5 | 62.5 |
| Moderate | 4 | 50 | Ongoing | 3 | 37.5 |
| Severe | 3 | 37.5 |  |  |  |
| Causality: relationship to study medical device | n | % | Causality: relationship to study procedure | n | % |
| Not Related | 7 | 87.5 | Not Related | 5 | 62.5 |
| Possible | 0 | 0 | Possible | 1 | 12.5 |
| Probable | 1 | 12.5 | Probable | 1 | 12.5 |
| Causal relationship (related) | 0 | 0 | Causal relationship (related) | 1 | 12.5 |

Table 21 Adverse events occurred (seriousness, severity, causal relationship with study medical device and study procedure, outcome of the event).

Table 22 Detailed description of adverse events occurred.

### 3.5.2 Device deficiency

No device deficiencies were recorded so far.

### 3.5.3 Device survivorship

As described in Table 22 of the Adverse Events sections, two subjects underwent revision surgery of Delta Revision cup and were thus withdrawn from the study. The serious adverse events that led to a revision surgery occurred to patient 101-9 and 101-14, 2 weeks and 6 weeks after surgery, respectively.

# SUMMARY AND OUTLOOK

Preliminary results suggest that the treatment with Delta Revision cup results in an improvement of clinical outcomes from preoperative to 2 year-follow-up. This can be observed both for HHS and OHS. No relevant radiographic issues have been identified from the radiographic analysis. Only 2 cases of revision have been recorded so far.