## Queries
Pt ID | FOLLOW UP | QUERY | ACTION | STATUS
101/04 | FU at 2 months / OXFORD | data are not inserted | please complete and validate | PENDING
101/4 | FU at 6 months/ OXFORD | data are not inserted | please complete and validate
101/5 | Adverse Event 20/10/2021 | is reported that the event is ANTICIPATED | expectedness should be ealuated only if the SAE is related to the procedure or to the device. In this case is not related, thus in this ase should be selected N/A option
101/10 | Patient data | stem is competitor and has not been removed. The statement has been inserted in the Adverse Event section | please, moe the statement and insert it in the Patient data NOTE section
101/13 | Adverse Event 12/04/2022 | the report is reported as "Initial" | Since it seems that the AE has been resolved during the surgery, please to update the report as "Final" and update the "Outcome of the Event" with the date.
101/15 | Patient data | stem is competitor and has not been removed | please insert the statement in the notes: Delta Revision Cup has been implanted in coupling with a competitor stem, because the stem was well osteointegrated. There was no need to remove to remove it.
101/1 to 101/15 | All FU | X-rays attachments are missing | to update the attachments on Limes | PENDING