## Queries_pending
Pt ID | FOLLOW UP | QUERY | ACTION | Sponsor comment | STATUS | SITE COMMENT | Status
101/1 | Adverse event 30/11/2021/involved study medical device information | As the medical device was not involved, please tick the "N/A" box in the sections "name of the medical device" and "components involved" | Please select "N/A" in the sections "name of the medical device involved" and "components invovled" |  | PENDING |  | SOLVED
101/1 | Adverse event 30/11/2021/event information | In the section "outcome of the event" the recovery date is missing | Please complete the section with end date of the adverse event |  | PENDING |  | PENDING
101/1 | Adverse event 30/11/2021/expectedness | The expectedness of the event is reported as "anticipated". However, evaluation of expectedness should be performed only when there is a causality with device or study procedure | Since in this case there is no relationship to the device or procedure, please select "N/A" in the "expectedness" section |  | PENDING
. | . | . | .
101/2 | FU 1 year | No data inserted in the eCRF | If the patient did not come back for the 1 year FU, please complete the section "missing follow up" with the missing reason. If the patient did come back for the 1 year FU visit, please complete the fields accordingly |  | PENDING
. | . | . | .
101/4 | Intraoperative | In the "stem" section, a Lima stem is reported. However, in the "notes" of the patient data section as well as in the "batch number/expiry date" section, it is stated that the stem is from competitor | Please revise these sections and correct them accordingly |  | PENDING
101/4 | FU 1 year/radiographic evaluation | Missing X-rays | Please upload missing x-rays once they have been properly anonymized |  | PENDING
101/4 | FU 1 year/radiographic evaluation | In the section "stem sizing", the evaluation is reported as "normal". However, in the previous FU visits (discharge, 2 months and 6 months FU) stem sizing is reported as "undersized" | Please review this field and adjust it if required |  | PENDING
. | . | . | .
101/6 | FU 1 year | No data inserted in the eCRF | If the patient did come back for the 1 year FU visit, please complete the fields accordingly. If the patient did not come back for the 1 year FU, please complete the section "missing follow up" with the missing reason. |  | PENDING
. | . | . | .
101/13 | FU 1 year | Visit date is the same for 6 months and 1 year FU | Please revise this field and adjust it accordingly |  | PENDING
101/13 | FU 1 year/radiographic evaluation | Visit date of x-rays is the same for 6 months and 1 year FU | Please revise this field and adjust it accordingly |  | PENDING
101/13 | FU 1 year/radiographic evaluation | The field "stem sizing" was not completed | Please complete the field "stem sizing" with the missing information |  | PENDING
. | . | . | .
101/16 | Patient data | The statement "Delta Revision Cup has been implanted in coupling with a competitor stem, because the stem was well osteointegrated. There was no need to remove it." should be added to the "Notes" field of the "patient data" section | Please move the sentence from the "notes" section of "batch number/expiry date" to the "notes" section of "patient data" |  | PENDING
101/16 | FU 6 months/radiographic evaluation | Missing X-rays | Please upload missing x-rays once they have been properly anonymized |  | PENDING
. | . | . | .
101/17 | Intraoperative | No data inserted in the eCRF | Please complete this section with missing data and validate the tab |  | PENDING
101/17 | Intraoperative/batch number/expiry data | No data inserted in the eCRF | Please complete this section with missing data and validate the tab |  | PENDING
101/17 | FU 6 months | No data inserted in the eCRF | Please complete this section with missing data, if available |  | PENDING
101/17 | Adverse event 24/04/2023 | Missing X-rays | Please upload missing x-rays once they have been properly anonymized |  | PENDING
101/17 | Adverse event 24/04/2023 | Causality/relationship to study device is reported as probable | Please revise these fields to evaluate if this evaluation is still the right one |  | PENDING
101/17 | Adverse event 24/04/2023 | Outcome of the event is still ongoing | Please check this section and update it with new information, if available |  | PENDING
. | . | . | .
101/18 | Patient data | The statement "Delta Revision Cup has been implanted in coupling with a competitor stem, because the stem was well osteointegrated. There was no need to remove it." should be added to the "Notes" field of the "patient data" section | Please move the sentence from the "notes" section of "surgery data" to the "notes" section of "patient data" |  | PENDING
101/18 | FU 2 months/radiographic evaluation | The section "osteolysis" is marked as "absent", however osteolysis is reported in 2 zones | Please revise this section and modofy the section "osteolysis" from "absent" to "present" |  | PENDING
. | . | . | .
101/21 | Patient data | The statement "Delta Revision Cup has been implanted in coupling with a competitor stem, because the stem was well osteointegrated. There was no need to remove it." should be added to the "Notes" field of the "patient data" section | Please move the sentence from the "notes" section of "batch number/expiry date" to the "notes" section of "patient data" |  | PENDING
101/21 | FU at discharge/radiographic evaluation | The section "osteolysis" is marked as "absent", however osteolysis is reported in 3 zones | Please revise this section and modofy the section "osteolysis" from "absent" to "present" |  | PENDING
101/21 | FU 2 months/radiographic evaluation | Missing X-rays | Please upload missing x-rays once they have been properly anonymized |  | PENDING
101/21 | FU 6 months/radiographic evaluation | Missing X-rays | Please upload missing x-rays once they have been properly anonymized |  | PENDING
101/21 | FU 6 months/radiographic evaluation | In the section "bone remodelling/stress shielding" no answer was reported | Please revise this section and complete it with "absent" or "present" |  | PENDING
. | . | . | .
101/22 | Patient data | The section "patient data" was not validated | Please validate this section |  | PENDING
101/22 | Patient data | The section "work status" was not completed | Please complete this section with the missing information, if available |  | PENDING
101/22 | FU at discharge/radiographic evaluation | The section "Osteolysis acetabular zone 3" was not completed | Please complete this section with the missing information, if available |  | PENDING
101/22 | FU 2 months | No data inserted in the eCRF | Please complete this section with the missing information, if available |  | PENDING
. | . | . | .
101/23 | All sections | Only the preoperative sections have been completed | Please complete intraoperative and postoperative sections with missing data, if available |  | PENDING
101/23 | Preoperative | In the section "previous hip treatments or surgeries on the affected side/if yes please specify" it is reported that there have 4 been 4 preivious operations on the affected hip | Please specifiy which operations occured to the affected hip |  | PENDING
101/23 | Preoperative/radiographic evaluation | Missing X-rays | Please upload missing x-rays once they have been properly anonymized |  | PENDING
. | . | . | .
101/24 | Patient data/concomitant medications | To the question "is the patient taking any concomitant medications?" the answer given is "-" | Since API and dosage have been specified in the following section, please change the answer from "-" to "yes" |  | PENDING
101/24 | Adverse event 12/07/2023 | In the section "invovled study medical device information" it can be deducted that the device involved refers to the previous prosthesis, this is not a Lima product | Please speficy next to "stem proxima" the word "(competitor)" to avoid misunderstaning |  | PENDING
| Adverse event 12/07/2023 | The section "expectedness" should be completed only there is a causality with medical device or study procedure |  |  | PENDING